## Slide 1: AEMEPS
27.09.2023

## Slide 2: Mouse test: A world of shite/leaks/EM interference etc?
It appears that the electric signal from the ultrasound is mixing with my applied signal. CHECK.
This means, I require electric isolation. Fill the transducer cone with mineral oil next time.
How is it that I am getting an area of mixing the size of the focus then – this part seems legit?
I am getting the trend where the neural signal is larger at lower frequencies… However, there is an open question as to how this is generated. The amplitude this time was so large, I think it was due to the e field of the transducer.
Do I see any spike datas?

## Slide 3
I suspect the previous time where I applied a dual acoustic frequency, is actually electric mixing. In that case – why is it focal with pressure?

## Slide 4: Analysis todos:
Plot the amplitude lfp w.r.t. phantom and difference frequency. Is it different in the mouse? I think it is… this might be the data from e105.
Look for spikes in this last experiment.
Look at the applied voltage signal and filter for the df. Can I see it and what amplitude is it at?
I can confirm that mineral oil is NOT conductive.

## Slide 5: Things I can actually do.
Fill the ultrasound cone with mineral oil. Can I still see the acoustoelectric effect if I use ae_calibrate_with_ae? Do I still see the PRF amplitude trend?
RF TI at 2MHz.

## Slide 6: Mouse:Acoustoelectric neuromodulation df = 1Hz. (e113 t1 file 12)
Time(s)
Volts(V)
Volts(V)
RF Ultrasound
Applied Voltage
Spikes (300-1000Hz)
Local Field Potential (0.5-300Hz)
Time(s)
Volts(V)
Volts(V)
Water filled US cone

## Slide 7
Mouse: Acoustoelectric neuromodulation df = 2Hz. (e113 t1 file21)
Volts(V)
Volts(V)
Applied Voltage Monitor
Note: This had the slightly leaky mineral oil layer.
Volts(V)
Volts(V)
Time(s)
Time(s)
RF Ultrasound
Spikes (300-1000Hz)
Local Field Potential (0.5-300Hz)

## Slide 8: Could the 2hz difference frequency be endogenous in its origin?
No signals applied (e113 t1 file 9 )
Applied Voltage Monitor
RF Ultrasound
Spikes (300-1000Hz)
Local Field Potential (0.5-300Hz)
RESULT: the 2Hz is NOT endogenous and there are no signs of it in the LFP or spikes,
though you can see some ketamine induced delta waves in the local field potential.
Volts(V)
Volts(V)
Time(s)
Time(s)

## Slide 9: Does any one applied signal alone produce a difference frequency of 2Hz?
1MPa Pressure alone (e113 t1 file 10 )
1V applied alone (e113 t1 file 11)
Local Field Potential (0.5-300Hz)
Spikes (300-1000Hz)
I unfortunately have EM noise in my rf amplifier which induces spikes over time…
but no 2Hz.
RESULT: No 2hz
These are endogenous delta rhythms induced from ketamine. (not at 2hz)
Spikes (300-1000Hz)
Local Field Potential (0.5-300Hz)
Volts(V)
Volts(V)
Volts(V)
Volts(V)

## Slide 10: Is the difference frequency in the applied signals at all? (e113 t1 file21)
Volts(V)
Volts(V)
Volts(V)
Local Field Potential (0.5-300Hz)
Applied voltage data filtered(0.5-300Hz)
Applied RF data filtered(0.5-300Hz)
These are all on the same scale, and it’s clear that there is no 2hz in the applied signals.
NOTE: THIS WHOLE SECTION IS INVALID because the signal isn’t captured in my daq resolution.
However, also – you would EXPECT the df voltage to be in the applied voltage because that’s how it works… so this is not a good test.
Due to the signal isolation, it’d be in the transformer output, but not the function generator.

## Slide 11: Is the difference frequency in the applied signals at all? (e113 t1 file21)
I could measure from the function generator for the voltage instead? This should not have the difference frequency in it because the transformer isolates. If I can show that the difference frequency amplitude is smaller than that measured in solution… still nothing because the transformer does 2x…

## Slide 12: Problem: the e field applied from the voltage source adjusts based on the US electrical PRF. (e113 t1 f6 - mouse) i.e. total harmonic distortion
df = 10. The applied voltage here, seems to clearly have the difference frequency within it, though it isn’t in the fft perhaps because I had a 40V daq range?
I can decrease the DAQ range to 2V and do a new phantom measurement to confirm if there is frequency mixing here. However, based on how the voltage source works, I do EXPECT to see the df anyway. I can also see the US PRF electrical signal in the FFT of the applied voltage source signal. See left.
Perhaps some of my f mixing amplitude is coming from non-linear mixing in the medium between the two electric signals?
This is a possibility – however, it’d be better to remove this option for non-linear signal distortion. This is possible with mineral oil.
Volts(V)
Time(s)
Frequency (Hz)
Volts(V)
Applied voltage source close up
Applied signal
Freeloader signal from the US

## Slide 13: Mineral Oil Solution (e114 mineral oil t1 f3)
Acoustoelectric difference frequency is still a good 2000 microvolts amplitude, making this a pure acoustoelectric interaction.
Frequency (Hz)
Frequency (Hz)
Applied voltage source close up shows the applied signal at 500100 massively dominating the US 500khz.
Applied voltage source
Time (s)
Volts(V)
Note: leaking mineral oil situation at the end. Messy. Also leaked on my mouse yesterday. It turns out that mineral oil dissolves parafilm.
Applied signal
Freeloader signal from the US is much smaller with mineral oil cone. This small amount may be coming from the cables etc. which means it would not be focal.
Applied signal
Electrically isolate the ultrasound cone from the mouse, such that only the acoustic wave gets through. I can try to do this with either F21 or mineral oil filled US cone. Currently, I do not know if either of these solutions blocks the pressure wave. Test both with the hydrophone.
Mineral Oil Cone(instead of water filled): I have massively reduced the signal getting into the voltage transformer.

## Slide 14: F21 looking at freeloader US signal. (e114 t2 f14)
Without thin piece of F21
Applied voltage source with no F21
With thin piece of F21
With F21
Do these have the same electrode positions?

## Slide 15: Does F21 work?
- set up the hydrophone and do a pressure measurement with and without the F21 in place. Does the amplitude significantly change? – Done.
- Do similar electrical characterization showing how much F21 attenuates electrically.
- it would be nice to have a difference frequency spatial map at 10Hz. This should be the same as the previous one I have with 8khz and 500khz.
- using ae_calibrate_with_prf.py it DOES look like the PRF*PRF amplitude is decreased with F21. Create a new map with this showing df/sf and carrier amplitudes as well, comparable with the previous map I made. Firstly, do a rough map without the F21, to show that I do get prf*prf focality… then do the nothing ness map.

## Slide 16: Does F21 let pressure pass through? (e115 t1)
Pressure (Pa)
Attenuation (%)= 2.7*F21 thickness(mm) + 9.2
F21 thickness (mm)
F21 thickness (mm)
Pressure Attenuation (%)
Pressure (Pa)= -12.3*F21 thickness(mm) + 373
These tests were done with the hydrophone, gel, F21 and a layer of glad wrap to protect the F21 from falling on the hydrophone tip. Sometimes there were small air bubbles trapped. I didn’t optimize the hydrophone position to be at the ultrasound focus, but otherwise the measurements are good and seemed reliable(use this data). I sent through a continuous wave 500kHz signal for 2 seconds with ramp on either side making a 4 second duration recording.

## Slide 17
Does F21 let electric fields pass through? (e115)
The voltage is the amplitude of the 500kHz as an electrical signal using PRF1020.
F21 significantly attenuates the voltage coming from the ultrasound transducer.
This is good, but doesn’t mean I can make super conclusive results? I could probably just use the thin piece of F21.
F21 thickness (mm)
F21 thickness (mm)
Volts (V)
Carrier Voltage Attenuation (%)

## Slide 18: DO MEPS again, this time with F21 comparison.
Also do an electrical test… in mouse, once I know the amplitudes.
Preamp non-linearity test too.

## Slide 19: Spatial PRF Test with F21. preamp gain =100

## Slide 20: Electrical signals only.
Apply the two signals at the amplitudes I see a neural response in saline, with the same preamp settings. i.e. 500khz prf and 500.001khz.
I can’t really do this so well.
Maybe try the other side of the preamp again?

## Slide 21: 10Hz focality plot, with mineral oil cone.
Theory, at a current of 500khz, the signal goes everywhere… do I get focality when I use a lower frequency e field?
The pressure signal e amplitude looks really high. Do a measure when connected, and not connected. It is surprisingly high both times. I think this is why I am not seeing focality.

## Slide 22: Water filled cone, focality plot, 10hz df.
The core problem is that I do not know if my preamp, or ultrasound is broken. Maybe both are fine.
A lot of the variance has to do with how I am measuring the df, and what is the distance between my measurement and reference probes is. It seems likely, that the US is delivering a large voltage and the mineral oil blocks it to some degree.
The core issue is that I cannot measure focality. I have past focality maps, but not one from today or when I measured the mouse, so I cannot be sure it is the ae effect. Basically – having a focal spot makes it the ae effect.
I can rule out acoustic/acoustic mixing with the hydrophone.

## Slide 23: PRF focality in mineral oil filled cone.
With carrier, 1020 and sum frequency info too. Same preamp settings as previous.

## Slide 24: Lastly: Preamp non-linearity test.

## Slide 25
Why did my file interpretation seem different from the files in the folder yesterday? Something weird happened? From the data analysis, mineral oil seemed like a great option.
Can I see 10Hz acoustoelectric focality using a mineral oil filled cone? (Previously I did this with the same two stim/measure electrodes..(e97_meps t3)) – would need to work on the daq ranges here.
What does the PRF look like using a mineral oil filled cone? Does it have focality?
Phantom test with higher resolution voltage monitor – do I see the difference frequency appear when using a water filled cone? I will then need to apply this signal separately -

## Slide 26: Could it be electrical mixing from the PRF of the applied ultrasound be mixing with the applied voltage? i.e. all EM mixing?
WE KNOW MIXING IS OCCURING as there is nothing at the difference frequency from the applied signals. (Wait: it is possible I am just below the noise floor of my voltage and rf monitors so I don’t know this for sure? )
However, we do not know if we are mixing the electric signal from the ultrasound with the applied electric signal, or the acoustic signal is mixing with the applied electric signal?
Fill cone with mineral oil as an electrical insulator.
RESULT: though it has no conductance as measured with multimeter, I can still see the electrical carrier signal around 4000 microvolts. When I add the shielding back to the cable it goes down to 1400 microvolts.
- do phantom test with it. PRF focality test. There is no focality. This is really interesting here… Need to do a big picture plot of this.
- ae_calibrate_with_ae 8khz… I am having trouble seeing the focality this way, I think because the current density is everywhere the pressure is. So what I need to do is have pressure, then the dual electrode has one being the current generator, the other the measurement electrode. The third is the reference.
I am currently not answering the question – is any pressure actually coming out with the mineral oil cone? To do this, I either need the hydrophone, or a different set up so I can capture the focus.
It turns out mineral oil dissolves parafilm film, so I had air bubbles…
- BIG NOTE: If I ramp the voltage to surpass the charge injection limit, I get a lot of non-linearities. This is certainly causing some of my problems. Also, the voltage from the transducer was entering into my applied signal – never good. The acoustoelectric neuromod paper should be done with the mineral oil in the cone.
- MEP dual signal.

## Slide 27: Alternate methods for determining whether it is caused by:
Electric signal mixing with acoustic signal.
Electric signal mixing with electric signal.
If this is the case, the stimulation area would not be focal. Show that it is focal. i.e. the PRF *PRF data that I already have?
Acoustic signal mixing with acoustic signal.
Show that this doesn’t occur through the PRF test I have already done.
What if the electric signal from the ultrasound is in solution, thus entering into the output of my signal generation? (well, that is why it is isolated…).. This… is also why it’d be better to shield the US electrically. Firstly though – prove that I am having this issue.
TODO: can I use F21 and measure pressure with the hydrophone?
Can I fill cone with mineral oil, and measure pressure with the hydrophone?
Do a focality map of the difference frequency generated at 10Hz.

## Slide 28: Next things to do:
Do a focality plot at the difference frequency of 10Hz, using just a water filled US cone 1MPa. This would just be of the difference frequency.
Test the F21 using the Hydrophone.
Prove that the electric signal from the US is altering the signal from the current generator. - signal analysis.
= = = = = = =
It would be better to do AE neuromodulation, when no(or at least very little) electrical signal from US reaches mouse.
Consider a new mouse, with a 3rd electrode for applying voltage, embedded in brain instead of outside.

## Slide 29: Amplitude trend with difference frequency
I need to repeat this and do it properly with comparison against a phantom. These are results from e113 t1(mouse with implanted electrodes), e107 t6(phantom). Though I apply voltage, where I apply it matters? Or does it?
Frequency (Hz)
Volts(V)

## Slide 30: Could the difference frequency be induced in the preamp?
Evidence 1: Since there are neural spikes, the difference frequency is induced before it enters the preamplifier.
Evidence 2: e105 t? (I will need to redo this once I have final amplitude data anyway, so I am holding off on the post processing plots as I know I can do it without mixing in the preamp) - will do it post oil test.

## Slide 31: Phantom Focality Plot
With and without F21? PRF and dual signal results?

## Slide 32: What current am I applying when I get a response? What is the E and J?
Look at e113 to find out.