## Slide 1: AEMEPS
27.09.2023

## Slide 2: Mouse test: A world of shite/leaks/EM interference etc?
It appears that the electric signal from the ultrasound is mixing with my applied signal. CHECK.
This means, I require electric isolation. Fill the transducer cone with mineral oil next time.
How is it that I am getting an area of mixing the size of the focus then – this part seems legit?
I am getting the trend where the neural signal is larger at lower frequencies… However, there is an open question as to how this is generated. The amplitude this time was so large, I think it was due to the e field of the transducer.
Do I see any spike datas?

## Slide 3
I suspect the previous time where I applied a dual acoustic frequency, is actually electric mixing. In that case – why is it focal with pressure?

## Slide 4: Mouse:Acoustoelectric neuromodulation df = 1Hz. (e113 t1 file 12)
Time(s)
Volts(V)
Volts(V)
RF Ultrasound
Applied Voltage
Spikes (300-1000Hz)
Local Field Potential (0.5-300Hz)
Time(s)
Volts(V)
Volts(V)
Water filled US cone

## Slide 5
Mouse: Acoustoelectric neuromodulation df = 2Hz. (e113 t1 file21)
Volts(V)
Volts(V)
Applied Voltage Monitor
Note: This had the slightly leaky mineral oil layer.
Volts(V)
Volts(V)
Time(s)
Time(s)
RF Ultrasound
Spikes (300-1000Hz)
Local Field Potential (0.5-300Hz)

## Slide 6: Could the 2hz difference frequency be endogenous in its origin?
No signals applied (e113 t1 file 9 )
Applied Voltage Monitor
RF Ultrasound
Spikes (300-1000Hz)
Local Field Potential (0.5-300Hz)
RESULT: the 2Hz is NOT endogenous and there are no signs of it in the LFP or spikes,
though you can see some ketamine induced delta waves in the local field potential.
Volts(V)
Volts(V)
Time(s)
Time(s)

## Slide 7: Does any one applied signal alone produce a difference frequency of 2Hz?
1MPa Pressure alone (e113 t1 file 10 )
1V applied alone (e113 t1 file 11)
Local Field Potential (0.5-300Hz)
Spikes (300-1000Hz)
I unfortunately have EM noise in my rf amplifier which induces spikes over time…
but no 2Hz.
RESULT: No 2hz
These are endogenous delta rhythms induced from ketamine. (not at 2hz)
Spikes (300-1000Hz)
Local Field Potential (0.5-300Hz)
Volts(V)
Volts(V)
Volts(V)
Volts(V)

## Slide 8: Is the difference frequency in the applied signals at all? (e113 t1 file21)
Volts(V)
Volts(V)
Volts(V)
Local Field Potential (0.5-300Hz)
Applied voltage data filtered(0.5-300Hz)
Applied RF data filtered(0.5-300Hz)
These are all on the same scale, and it’s clear that there is no 2hz in the applied signals.
NOTE: THIS WHOLE SECTION IS INVALID because the signal isn’t captured in my daq resolution.
However, also – you would EXPECT the df voltage to be in the applied voltage because that’s how it works… so this is not a good test.
Due to the signal isolation, it’d be in the transformer output, but not the function generator.

## Slide 9: Is the difference frequency in the applied signals at all? (e113 t1 file21)
I could measure from the function generator for the voltage instead? This should not have the difference frequency in it because the transformer isolates. If I can show that the difference frequency amplitude is smaller than that measured in solution… still nothing because the transformer does 2x…

## Slide 10: Problem: the e field applied from the voltage source adjusts based on the US electrical PRF. (e113 t1 f6 - mouse) i.e. total harmonic distortion
df = 10. The applied voltage here, seems to clearly have the difference frequency within it, though it isn’t in the fft perhaps because I had a 40V daq range?
I can decrease the DAQ range to 2V and do a new phantom measurement to confirm if there is frequency mixing here. However, based on how the voltage source works, I do EXPECT to see the df anyway. I can also see the US PRF electrical signal in the FFT of the applied voltage source signal. See left.
Perhaps some of my f mixing amplitude is coming from non-linear mixing in the medium between the two electric signals?
This is a possibility – however, it’d be better to remove this option for non-linear signal distortion. This is possible with mineral oil.
Volts(V)
Time(s)
Frequency (Hz)
Volts(V)
Applied voltage source close up
Applied signal
Freeloader signal from the US

## Slide 11: Mineral Oil Solution (e114 mineral oil t1 f3)
Acoustoelectric difference frequency is still a good 2000 microvolts amplitude, making this a pure acoustoelectric interaction.
Frequency (Hz)
Frequency (Hz)
Applied voltage source close up shows the applied signal at 500100 massively dominating the US 500khz.
Applied voltage source
Time (s)
Volts(V)
Note: leaking mineral oil situation at the end. Messy. Also leaked on my mouse yesterday. It turns out that mineral oil dissolves parafilm.
Applied signal
Freeloader signal from the US is much smaller with mineral oil cone. This small amount may be coming from the cables etc. which means it would not be focal.
Applied signal
Electrically isolate the ultrasound cone from the mouse, such that only the acoustic wave gets through. I can try to do this with either F21 or mineral oil filled US cone. Currently, I do not know if either of these solutions blocks the pressure wave. Test both with the hydrophone.
Mineral Oil Cone(instead of water filled): I have massively reduced the signal getting into the voltage transformer.

## Slide 12: F21 looking at freeloader US signal. (e114 t2 f14)
Without thin piece of F21
Applied voltage source with no F21
With thin piece of F21
With F21
Do these have the same electrode positions?

## Slide 13: Does F21 let pressure pass through? (e115 t1)
Pressure (Pa)
Attenuation (%)= 2.7*F21 thickness(mm) + 9.2
F21 thickness (mm)
F21 thickness (mm)
Pressure Attenuation (%)
Pressure (Pa)= -12.3*F21 thickness(mm) + 373
These tests were done with the hydrophone, gel, F21 and a layer of plastic wrap to protect the F21 from falling on the hydrophone tip. Sometimes there were small air bubbles trapped. I didn’t optimize the hydrophone position to be at the ultrasound focus, but otherwise the measurements are good and seemed reliable(use this data). I sent through a continuous wave 500kHz signal for 2 seconds with ramp on either side making a 4 second duration recording.

## Slide 14
Does F21 let electric fields pass through? (e115)
The voltage is the amplitude of the 500kHz as an electrical signal using PRF1020 @ 500kHz.
F21 significantly attenuates the voltage coming from the ultrasound transducer. This is a good addition, but I still need the electrical only mixing test.
I could probably just use the thin piece of F21 which is easiest to use.
F21 thickness (mm)
F21 thickness (mm)
Volts (V)
Carrier Voltage Attenuation (%)

## Slide 15: Does F21 work? Yes – reasonably well.
- set up the hydrophone and do a pressure measurement with and without the F21 in place. Does the amplitude significantly change? – Done.
- Do similar electrical characterization showing how much F21 attenuates electrically. – Done.
- it would be nice to have a difference frequency spatial map at 10Hz. This should be the same as the previous one I have with 8khz and 500khz. – TODO. I am having issues seeing any focality lately, since the transducer started emitting these large artefact voltages.
- using ae_calibrate_with_prf.py it DOES look like the PRF*PRF amplitude is decreased with F21. Create a new map with this showing df/sf and carrier amplitudes as well, comparable with the previous map I made. Firstly, do a rough map without the F21, to show that I do get prf*prf focality… then do the nothing ness map. – Done. I need to look into how the position of my electrodes effects the background signal amplitude. As I got a constant 30 microvolts… (non-focal).
-I also got continuous, dual and PRF signal repeats to compare with hydrophone data for possible first section of mechanism of US paper.
Summary: Since F21 works, I am in a good place to revise mouse experiments. This time using F21, then no F21 as a comparison metric. I should also add in an electrical non-linearity test.

## Slide 16: Spatial PRF 1020 test with F21. preamp gain =100
TODO: look at the raw data signal before and after placement of F21.
0
32
There has been a weird change in how the US transducer functions whereby it emits a much larger EM noise for the same amplitude input signal from the RF amplifier. Should I send it back to Precision acoustics for checking? I want it to be up to spec. Note: There is a frayed cord area near the ultrasound base, which I have tried to fix with rubber tape… it is a likely weak point. TODO: Discuss with Nir.
I would have expected the result to be no focality here, which I get… which is good… BUT I see the baseline has changed to about 32 microvolts for dF. Where before I got noise floor around-2microvolt background. What is going on? Is this related to the EM noise increase from the transsducer? (I have confirmed with hydrophone I still get pressure out, and the RF amplifier is behaving properly. I think maybe the live US wire is touching GND somewhere where it didn’t before or something like that.)

## Slide 17: Phantom electrical equivalent two-tone test w F21 (e115 t3) – preamp gain = 1 (no filters)
Volts(V)
Frequency (Hz)
Volts(V)
df = 10Hz, US applied(1MPa), At measurement electrodes. File 12
Volts(V)
Volts(V)
Volts(V)
Volts(V)
Frequency (Hz)
Time(s)
Time(s)
Electrical two-tone test with signals at equivalent (and slightly higher) amplitudes at measurement electrodes. File 7
0.5-300Hz filter LFP filter
10Hz df clearly visible
0.5-300Hz filter LFP filter
No 10Hz df
V signal
US
V signal
V2 signal

## Slide 18: Phantom Two-tone difference frequency comparison.
Volts(V)
P = 1.4e-7
***
This is at my noise floor for this DAQ range setting.
TODO: Put no spikes from phantom time series.
Compare shapes of LFPs between them.
More comparison on US alone, vs V alone. Any spiking?
Do similar non-linearity test at the spiking frequency range. 300-1000hz. Prove it is not artefactually leaking from either stim alone.
Create a control response, i.e. a step function of e stim, i.e. 1Hz – do I see the same spiking neural response? (i.e. compare the neural response with acoustoelectric spiking with this). i.e. does 1hz e stim look the same.
Is there delay between applied stim and lfp? Is there a delay in the spiking too?

## Slide 19: Mouse with F21 (e113 t2)
Df = 2Hz, spikes, LFP
Df = 5Hz, spikes, LFP
Volts(V)
Volts(V)
1MPa and 2V output in RF resonator configuration
Spikes (300-1000Hz)
Time(s)
Time(s)
Local Field Potential (0.5-300Hz)

## Slide 20: 1hz df.

## Slide 21: Does any one applied signal alone produce a difference frequency of 2Hz?
1MPa Pressure alone (e113 t1 file 10 )
1V applied alone (e113 t1 file 11)
Local Field Potential (0.5-300Hz)
Spikes (300-1000Hz)
I unfortunately have EM noise in my rf amplifier which induces spikes over time…
but no 2Hz.
RESULT: No 2hz
These are endogenous delta rhythms induced from ketamine. (not at 2hz)
Spikes (300-1000Hz)
Local Field Potential (0.5-300Hz)
Volts(V)
Volts(V)
Volts(V)
Volts(V)

## Slide 22: Could the 2hz difference frequency be endogenous in its origin?
No signals applied (e113 t1 file 9 )
Applied Voltage Monitor
RF Ultrasound
Spikes (300-1000Hz)
Local Field Potential (0.5-300Hz)
RESULT: the 2Hz is NOT endogenous and there are no signs of it in the LFP or spikes,
though you can see some ketamine induced delta waves in the local field potential.
Volts(V)
Volts(V)
Time(s)
Time(s)

## Slide 23
Electrical two-tone test vs acoustoelectric mixing in vivo (e113 t2)
Electrical two-tone test with signals at equivalent (and slightly higher) amplitudes at measurement electrodes. File 16
Volts(V)
Volts(V)
Volts(V)
Frequency (Hz)
Time(s)
0.5-300Hz filter LFP filter
V signal
V2 signal
Volts(V)
Frequency (Hz)
Volts(V)
df = 10Hz, US applied(1MPa), At measurement electrodes. File 18
Volts(V)
Time(s)
0.5-300Hz filter LFP filter
V signal
US
10Hz df clearly visible
No 10Hz df

## Slide 24
In vivo Two-tone difference frequency comparison.
P = 2.9e-8
***
Volts(V)

## Slide 25: Df amplitude with frequency comparison in phantom/mouse
TODO: look at how the applied signal amplitudes change w.r.t. df.
Mouse with F21 (e113 t2)
Volts(V)
Note:
There is a lot of variance based on unknowns… for instance I suspect the volume of the conductive gel in the phantom and the volume of the mouse make a difference, as does how close the mouse and/or phantom is to GND plane. These plots are currently not making any statements.
Frequency (Hz)
Volts(V)
Mouse without F21 (e113 t1)

## Slide 26: Is this sufficient for a proof of acoustoelectric neuromodulation?
Cons – it relies heavily on spiking data as in vivo evidence. Is there another source of neural evidence?
Behavioral data is a bit hard to get… but would be ideal. Maybe I should really focus on getting this?
It seems pretty convincing otherwise… lots of artefact tests etc.
The measurement I am making is a differential measurement which means I am rejecting common mode noise. However, I think the common mode amplitude is quite important to the acoustoelectric effect and how big it is.
*I do think the US transducer has a weird issue where its EM profile has changed for the worse recently meaning I have much higher EM noise such that the GNDed amplitude of the 500khz is super huge, but I don’t think this confounds the tests I have done, for instance the two-tone non-linearity test seems like it should be a fine test against this.

## Slide 27
TODO: Put no spikes from phantom time series.
Compare shapes of LFPs between phantom and mouse. Delay there?
Is there delay between applied stim and lfp? Is there a delay in the spiking too?
More comparison on US alone, vs V alone. Any spiking? – look close at amplitudes.
Do similar non-linearity test at the spiking frequency range. 300-1000hz. Prove it is not artefactually leaking from either stim alone.
Create a control response, i.e. a step function of e stim, i.e. 1Hz – do I see the same spiking neural response? (i.e. compare the neural response with acoustoelectric spiking with this). i.e. does 1hz e stim look the same.

## Slide 28: 10Hz focality plot, with mineral oil cone.
Theory, at a current of 500khz, the signal goes everywhere… do I get focality when I use a lower frequency e field?
The pressure signal e amplitude looks really high. Do a measure when connected, and not connected. It is surprisingly high both times. I think this is why I am not seeing focality.

## Slide 29: Water filled cone, focality plot, 10hz df.
The core problem is that I do not know if my preamp, or ultrasound is broken. Maybe both are fine.
A lot of the variance has to do with how I am measuring the df, and what is the distance between my measurement and reference probes is. It seems likely, that the US is delivering a large voltage and the mineral oil blocks it to some degree.
The core issue is that I cannot measure focality. I have past focality maps, but not one from today or when I measured the mouse, so I cannot be sure it is the ae effect. Basically – having a focal spot makes it the ae effect.
I can rule out acoustic/acoustic mixing with the hydrophone.

## Slide 30
Why did my file interpretation seem different from the files in the folder yesterday? Something weird happened? From the data analysis, mineral oil seemed like a great option.
Can I see 10Hz acoustoelectric focality using a mineral oil filled cone? (Previously I did this with the same two stim/measure electrodes..(e97_meps t3)) – would need to work on the daq ranges here.
What does the PRF look like using a mineral oil filled cone? Does it have focality?
Phantom test with higher resolution voltage monitor – do I see the difference frequency appear when using a water filled cone? I will then need to apply this signal separately -

## Slide 31: Could it be electrical mixing from the PRF of the applied ultrasound be mixing with the applied voltage? i.e. all EM mixing?
WE KNOW MIXING IS OCCURING as there is nothing at the difference frequency from the applied signals. (Wait: it is possible I am just below the noise floor of my voltage and rf monitors so I don’t know this for sure? )
However, we do not know if we are mixing the electric signal from the ultrasound with the applied electric signal, or the acoustic signal is mixing with the applied electric signal?
Fill cone with mineral oil as an electrical insulator.
RESULT: though it has no conductance as measured with multimeter, I can still see the electrical carrier signal around 4000 microvolts. When I add the shielding back to the cable it goes down to 1400 microvolts.
- do phantom test with it. PRF focality test. There is no focality. This is really interesting here… Need to do a big picture plot of this.
- ae_calibrate_with_ae 8khz… I am having trouble seeing the focality this way, I think because the current density is everywhere the pressure is. So what I need to do is have pressure, then the dual electrode has one being the current generator, the other the measurement electrode. The third is the reference.
I am currently not answering the question – is any pressure actually coming out with the mineral oil cone? To do this, I either need the hydrophone, or a different set up so I can capture the focus.
It turns out mineral oil dissolves parafilm film, so I had air bubbles…
- BIG NOTE: If I ramp the voltage to surpass the charge injection limit, I get a lot of non-linearities. This is certainly causing some of my problems. Also, the voltage from the transducer was entering into my applied signal – never good. The acoustoelectric neuromod paper should be done with the mineral oil in the cone.
- MEP dual signal.

## Slide 32: Alternate methods for determining whether it is caused by:
Electric signal mixing with acoustic signal.
Electric signal mixing with electric signal.
If this is the case, the stimulation area would not be focal. Show that it is focal. i.e. the PRF *PRF data that I already have?
Acoustic signal mixing with acoustic signal.
Show that this doesn’t occur through the PRF test I have already done.
What if the electric signal from the ultrasound is in solution, thus entering into the output of my signal generation? (well, that is why it is isolated…).. This… is also why it’d be better to shield the US electrically. Firstly though – prove that I am having this issue.
TODO: can I use F21 and measure pressure with the hydrophone?
Can I fill cone with mineral oil, and measure pressure with the hydrophone?
Do a focality map of the difference frequency generated at 10Hz.

## Slide 33: Next things to do:
Do a focality plot at the difference frequency of 10Hz, using just a water filled US cone 1MPa. This would just be of the difference frequency.
Test the F21 using the Hydrophone.
Prove that the electric signal from the US is altering the signal from the current generator. - signal analysis.
= = = = = = =
It would be better to do AE neuromodulation, when no(or at least very little) electrical signal from US reaches mouse.
Consider a new mouse, with a 3rd electrode for applying voltage, embedded in brain instead of outside.

## Slide 34: Amplitude trend with difference frequency
I need to repeat this and do it properly with comparison against a phantom. These are results from e113 t1(mouse with implanted electrodes), e107 t6(phantom). Though I apply voltage, where I apply it matters? Or does it?
Frequency (Hz)
Volts(V)

## Slide 35: Could the difference frequency be induced in the preamp?
Evidence 1: Since there are neural spikes, the difference frequency is induced before it enters the preamplifier.
Evidence 2: e105 t? (I will need to redo this once I have final amplitude data anyway, so I am holding off on the post processing plots as I know I can do it without mixing in the preamp) - will do it post oil test.

## Slide 36: Phantom Focality Plot
With and without F21? PRF and dual signal results?

## Slide 37: What current am I applying when I get a response? What is the E and J?
Look at e113 to find out.